## Slide 1
MOD·03
TEST MODULE
Test &
Measurement
The instruments that turn guessing into diagnosis.
MULTIMETER
SCOPE
ESR METER
SUPPLY
SIGNAL

## Slide 2
MOD·03
TEST · OVERVIEW
The Five Instruments on Every Bench
Each one answers a different question about a faulty board.
Multimeter
Voltage, resistance, continuity
Every job
Oscilloscope
Shows a signal changing over time
Waveform
ESR Meter
Tests capacitors in-circuit
In-circuit
Bench Supply
A safe, adjustable power source
Adjustable
Signal Generator
Injects a known signal to trace faults
Inject

## Slide 3
MOD·03
TEST · FUNDAMENTALS
What Is a Measurement?
A number only means something next to what it should be.
In Plain Terms
A measurement compares what a circuit is actually doing against what it should be doing — that gap is the fault.
Real-World Picture
Like a thermometer on a patient — the number alone means nothing until you compare it to what's normal.
Key Facts
Every reading needs a reference: a datasheet value, a known-good board, or a calculated expectation.
No reference, no diagnosis

## Slide 4
MOD·03
TEST · FUNDAMENTALS
Live vs Offline Testing
The same component, tested two very different ways.
1
Offline testing
Component removed or isolated — its true, standalone value
2
In-circuit testing
Measured on the board — faster, but nearby parts can skew it
3
Live testing
Power on, circuit running — the only way to catch some faults
4
Choosing which to use
Start offline when you can; go live only when the fault demands it
In-circuit readings can lie — know when to isolate

## Slide 5
MOD·03
TEST · FUNDAMENTALS
Reading It Right
A wrong setting gives a confident, wrong answer.
HOW TO GET AN ACCURATE READING
1
Match the range and function to what you expect
2
Let the reading settle before recording it
3
Zero or calibrate the meter first, if it needs it
COMMON MISTAKES
!
Reading in-circuit and assuming it’s the true value
!
Wrong probe sockets for the function selected
!
Trusting a single reading with no reference point
One good reading beats three guesses

## Slide 6
MOD·03
TEST · EXAMPLE
Real Example: Multimeter
A dead phone charger, diagnosed in under a minute.
The Scenario
A phone charger produces no output at all. The customer says it just stopped working overnight.
What You Do
Set the meter to DC voltage and probe the output pins while the charger is plugged into the wall.
What It Tells You
A 0V reading confirms the fault sits upstream, inside the charger — not the cable or the phone.
0V out = fault confirmed

## Slide 7
MOD·03
TEST · FUNDAMENTALS
Multimeter Types & Categories
Not all multimeters are built for the same job.
1
Analog vs digital
Digital gives a precise number; analog shows a trend as it needles across the scale
2
Auto-ranging
Picks the right range automatically — fewer mistakes, slightly slower to settle
3
True-RMS
Reads AC accurately even on non-sine waveforms — standard on modern meters
4
CAT rating
Matches the meter to the energy level of the circuit it’s rated to test safely
The CAT rating is a safety spec, not a marketing number

## Slide 8
MOD·03
TEST · FUNDAMENTALS
Reading Multimeter Markings
The label tells you what the meter can — and can’t — safely do.
1
CAT rating & voltage
e.g. CAT III 600V — the environment and energy level it’s rated for
2
Accuracy (%)
How close a reading is to the true value — tighter for lab-grade meters
3
Fuse rating
The current-mode fuse rating — replace only with the exact spec
4
Display count
e.g. 6000-count — how fine a resolution the display can show
A blown fuse often looks just like a dead meter

## Slide 9
MOD·03
TEST · FUNDAMENTALS
Common Multimeter Failures
Before you blame the circuit, check the meter itself.
STILL RELIABLE
1
Battery indicator steady, display crisp
2
Leads flex without any flickering reading
3
Continuity beep is sharp and consistent
NEEDS ATTENTION
!
Reads nothing in current mode — check the fuse first
!
Reading drifts or flickers with the leads still
!
Cracked or intermittent test leads
A dead fuse is more common than a dead meter

## Slide 10
MOD·03
TEST · OSCILLOSCOPE
What Is an Oscilloscope?
It shows you what a meter can only describe in numbers.
In Plain Terms
An oscilloscope draws a graph of voltage over time — so you can see a signal’s shape, not just its value.
Real-World Picture
Like a heart-rate monitor — the shape of the trace tells you as much as the number does.
Key Facts
A multimeter gives one number; a scope shows you if that number is steady, pulsing, or corrupted by noise.
Seeing the shape finds what a meter can’t

## Slide 11
MOD·03
TEST · OSCILLOSCOPE
Oscilloscope Specifications That Matter
Three numbers decide what a scope can actually catch.
Bandwidth
The highest frequency the scope can measure accurately — buy above what you need
Sample rate
How many points per second it captures — too low and fast events are missed entirely
Channels
Two or more inputs let you compare two signals against each other at once
Bandwidth you don’t have is a fault you’ll never see

## Slide 12
MOD·03
TEST · OSCILLOSCOPE
What It Shows You
Four things a waveform tells you at a glance.
1
Shape
Sine, square, or a mess — the shape reveals the signal’s health
2
Amplitude (voltage)
How tall the wave is — is it reaching the level it should?
3
Frequency & period
How fast the signal repeats — timing faults show up here
4
Noise & ripple
Unwanted wobble riding on a signal that should be clean
A clean signal is a smooth, predictable trace

## Slide 13
MOD·03
TEST · OSCILLOSCOPE
Types of Oscilloscope Probes
The probe shapes the signal before the scope even sees it.
1
1x probe
No attenuation — better for very small signals, but loads the circuit more
2
10x probe
Attenuates the signal — the everyday default for most bench work
3
Ground clip
Always connect it to the circuit’s common, or the trace becomes meaningless
4
Probe compensation
A quick square-wave check ensures the probe isn’t distorting the signal
An uncompensated probe lies to you convincingly

## Slide 14
MOD·03
TEST · OSCILLOSCOPE
Set It Up & Read It
The right settings turn noise into a readable trace.
HOW TO SET UP
1
Set volts/division to fit the expected signal size
2
Set time/division to fit the expected frequency
3
Trigger on a stable point in the waveform
COMMON MISTAKES
!
Probe not grounded to the circuit’s common
!
Scale set too coarse to see the real detail
!
Misreading the grid because the scale wasn’t noted
Trigger first — a still trace is a readable trace

## Slide 15
MOD·03
TEST · EXAMPLE
Real Example: Oscilloscope
A flickering LED strip, traced to a noisy driver.
The Scenario
An LED strip flickers randomly, but the multimeter shows a perfectly steady average voltage.
What You Do
Probe the driver’s output with the oscilloscope instead of relying on the multimeter alone.
What It Tells You
The trace reveals voltage ripple the multimeter had quietly averaged away — the real fault.
Ripple hides inside an average reading

## Slide 16
MOD·03
TEST · SPECIALIST METERS
ESR Meter — Testing Capacitors In-Circuit
Finds a weak capacitor without desoldering it first.
In Plain Terms
ESR (Equivalent Series Resistance) measures a capacitor’s internal resistance — high ESR means a worn-out part.
Real-World Picture
Like checking a battery’s health, not just its voltage — a capacitor can read the right value and still be bad.
Key Facts
ESR testing works in-circuit, on a discharged board — no desoldering needed to catch most bad capacitors.
Right capacitance, high ESR = still faulty

## Slide 17
MOD·03
TEST · SPECIALIST METERS
More Specialist Tools
Reach for these when the multimeter isn’t enough.
1
LCR meter
Precise inductance, capacitance, and resistance, out of circuit
2
Insulation tester
Checks that a winding or cable isn’t leaking current where it shouldn’t
3
Clamp meter
Reads current without breaking the circuit open
4
Thermal camera
Spots a component running hotter than its neighbours
Specialist tools save hours on stubborn faults

## Slide 18
MOD·03
TEST · EXAMPLE
Real Example: ESR Meter
A power supply that boots, then randomly resets.
The Scenario
A PC power supply turns on fine but resets under load. Every capacitor measures the correct value.
What You Do
Test the same capacitors in-circuit with an ESR meter instead of a plain capacitance meter.
What It Tells You
A high ESR reading reveals a worn capacitor that the capacitance test alone couldn’t catch.
Right value, high ESR = still bad

## Slide 19
MOD·03
TEST · POWER & SIGNAL
The Bench Power Supply
A safe, adjustable stand-in for the board’s real power.
In Plain Terms
A bench supply provides a clean, adjustable voltage with a current limit — so a fault can’t do further damage.
Real-World Picture
Like testing a car engine on a stand before it goes back on the road — controlled, and safe to fail.
Key Facts
Set the current limit low first — it protects the board (and you) if a short is still present.
Current limit first, voltage second

## Slide 20
MOD·03
TEST · EXAMPLE
Real Example: Bench Power Supply
Testing a repaired board without risking it twice.
The Scenario
A board just had a shorted component replaced. Powering it straight from the wall is still risky.
What You Do
Power it from a bench supply with the current limit set low before switching it on.
What It Tells You
If a short remains, the limit trips safely — no second burnt component to replace.
The limit trips before the damage does

## Slide 21
MOD·03
TEST · POWER & SIGNAL
Function Generator & Signal Injection
Feed in a known signal to trace where it gets lost.
1
Sine & square waves
Standard test signals with known shape, amplitude, and frequency
2
Injecting a signal
Feed it in at a known point and follow it through the circuit
3
Tracing a fault
The signal disappears exactly where the fault sits
4
Safe limits
Match amplitude to the circuit’s rating — too much damages good parts
Follow the signal until it goes quiet

## Slide 22
MOD·03
TEST · EXAMPLE
Real Example: Signal Generator
Tracing a dead audio channel through an amplifier.
The Scenario
One channel of a stereo amplifier is silent, and the fault could be anywhere along the signal chain.
What You Do
Inject a test tone at the input and follow it stage by stage with a scope or meter.
What It Tells You
The tone disappears exactly at the faulty stage — no more guessing which part failed.
The signal stops exactly at the fault

## Slide 23
MOD·03
TEST · SAFETY & METHOD
Choosing the Right Test Tool
The fastest diagnosis starts with the right instrument.
1
Suspect a dead supply?
Start with the multimeter — voltage first, always
2
Suspect noise or timing?
Reach for the oscilloscope — numbers alone won’t show it
3
Board resets under load?
Check capacitors in-circuit with an ESR meter
4
Need to isolate a stage?
Inject a signal and trace it with the function generator
The right tool turns a guess into a confirmed fault

## Slide 24
MOD·03
TEST · SAFETY & METHOD
PCB Identification: Test Points
Boards label their own test points — if you know what to look for.
1
TP1, TP2… labels
Small silkscreen labels marking a designer-intended measurement point
2
GND symbol
Marks a reliable ground reference for your black probe
3
Rail labels (VCC, 5V, 3V3)
Tells you the expected voltage before you even measure it
4
Unlabelled test pads
Bare copper pads with no silkscreen — check the schematic if one exists
A labelled board is the designer helping you troubleshoot

## Slide 25
MOD·03
TEST · SAFETY & METHOD
Test Safely, Every Time
The same discipline from Module 1, applied to every instrument.
HOW TO TEST SAFELY
1
Match the meter’s CAT rating to the circuit
2
Isolate and discharge before offline testing
3
Start on the highest safe range, then narrow down
WARNING SIGNS
!
Meter reading is unstable or won’t settle
!
Smell of burning or a warm test lead
!
Reading far outside any reasonable expected value
The right instrument, used the right way

## Slide 26
MOD·03
TEST · PRACTICAL
Practical Exercise: Build a Test Log
Measure a known-good practice board and record what you find.
✓
Set the multimeter to DC voltage and record the supply rail
✓
Check continuity on five marked test points
✓
Probe one signal line with the oscilloscope and sketch the trace
✓
Test one capacitor in-circuit with the ESR meter
✓
Compare every reading to the datasheet or a known-good board
✓
Hand your completed log to your instructor
A logged reading is worth more than a remembered one

## Slide 27
MOD·03
TEST · REPAIR EXERCISE
Repair Exercise: Diagnose With the Right Tool
A board resets under load — and the wrong tool nearly hid the fault.
The Scenario
A board powers on fine but resets under load. A multimeter shows every capacitor at the right value.
Wrong Tool Used
Sticking with capacitance mode alone missed a capacitor with high ESR but a correct value.
Right Approach
Switching to an ESR meter, in-circuit, revealed the one capacitor a value check couldn’t catch.
The right value doesn’t always mean the right part

## Slide 28
MOD·03
TEST · QUIZ
Quick Check
Talk through these before moving on to Module 4.
1
Question 1
Why can an in-circuit reading be misleading compared to an offline one?
2
Question 2
What does an oscilloscope show you that a multimeter never can?
3
Question 3
Why might a capacitor read the correct value and still be faulty?
4
Question 4
Why set a bench supply’s current limit low before powering a repaired board?
Discuss as a group before checking the answer key

## Slide 29
MOD·03
TEST · QUICK REFERENCE
Quick Instrument Reference
One line per instrument — what it measures, and what to watch for.
INSTRUMENT
MEASURES
TYPICAL SETTING
WATCH FOR
Multimeter
V · Ω · continuity · A
CAT II+
Wrong function selected
Oscilloscope
Waveform over time
Auto-trigger first
Ungrounded probe
ESR meter
Capacitor health
In-circuit, discharged
Board still charged
Bench supply
Adjustable V · A limit
Low current limit first
Limit set too high
Signal generator
Test sine · square wave
Match circuit rating
Amplitude too high

## Slide 30
MOD·03
MODULE WRAP-UP
You Can Now…
By the end of this module, every recruit should be able to:
✓
Choose live, in-circuit, or offline testing for a fault
✓
Read a waveform on an oscilloscope with confidence
✓
Test a capacitor’s health in-circuit with an ESR meter
✓
Use a bench supply and signal generator safely
✓
Trust a reading because you know what it should be
!
GOLDEN RULE
A measurement means nothing without something to compare it to.